## Slide 1: Biomedical Imaging X-Ray Imaging (II)
Sebastian Kozerke
Institute for Biomedical Engineering, University and ETH Zurich, Switzerland

## Slide 2: Today’s Learning Objectives
Describe basic X-ray detection technologies
Relate image contrast to photon scattering
Explain encoding and decoding of projection data
Derive filtered back projection principle
Implement CT encoding and decoding (Exercise)

## Slide 3: Resolution of imaging procedures
Spatial resolution
    Blurring of object points when projected into image
Contrast
    Difference in linear attenuation coefficients between tissues
Temporal resolution
    Sample rate of temporal recording
Chapter 1.5

## Slide 4: Resolution – Contrast
bone
muscle
fat
E [keV]
μ [cm-1]
Notes: Blackboard
-> scetch object with bone and soft tissue
-> illustrate weighted average of  in case of heterogeneous composition along d

## Slide 5: Resolution – Contrast
Large contrast between bone and tissue/air
Very little contrast among different tissues
Notes: Blackboard
-> I/I0=exp(-d) -> display  ~ -log(I/I0)
-> calculate maximum contrast between air and average biological tissue with d = 20 cm (and =0.35 cm-1 w/o bone or air): I/I0 = exp(-d) = exp(-7) = 1/1000
-> calculate maximum contrast between air and bone tissue with d = 20 cm (and =0.7 cm-1): I/I0 = exp(-d) = exp(-14) = 8/10000000

## Slide 6: Resolution of imaging procedures
Spatial resolution
    Blurring of object points when projected into image
Contrast
    Difference in linear attenuation coefficients between tissues
Temporal resolution
    Sample rate of temporal recording

## Slide 7: Resolution – Temporal resolution
Notes: Blackboard
-> scetch ventricular volume as function of time + volume curve
-> scetch Fourier pair of volume curve: Δf = 1/T = 1 Hz

## Slide 8
Detector technology
Chapter 2.6-2.7

## Slide 9: Projection imaging – Photographic film
Double-sided film with intensifying screens
film*)
carrier
*) Silver halide (AgX): compound formed between silver and silver bromide (AgBr) or silver chloride (AgCl) or silver iodide (AgI).
fluorescent screen #1 **)
fluorescent screen #2 **)
**) Fluorescence increases efficiency about 5-fold but decreases resolution
Notes: Blackboard
-> scetch thin and thick layer of AgX emulsion -> explain oblique incident radiation + scatter -> trade-off between layer thickness and absorption (-> Beer-Lambert)
-> scetch S = log(Opacity) as function of log(Exposure) -> linear range ΔS = 2.5

## Slide 10: Projection imaging – Anti scatter grids
Film with anti-scatter grid
lead lamellae
Notes: Blackboard
-> 1) scetch primary image (without scatter)
-> 2) scetch primary + scatter image (same image profile as in 1) but background signal larger)

## Slide 11: Projection imaging – Image intensifier
scintillating screen + photocathode
e-
e-
e-
e-
e-
UDC
photoanode + digital
output screen
incident radiation
Notes: Notes
An X-ray image intensifier (XRII) refers to a special image intensifier device used in medical imaging involving x rays. It consists of an input window, input phosphor, photocathode, vacuum and electron optics, output phosphor and output window. It allows for lower x-ray doses to be used on patients by magnifying the intensity produced in the output image, enabling the viewer to easily see the structure of the object being imaged.

## Slide 12: Projection imaging – Image intensifier
Notes: Blackboard
-> scetch:
Input screen (CsI) -> Luminescence –> x-ray -> photon
Photokathode -> photons -> electrons get accelerated

## Slide 13: Projection imaging – Summary
102 104 106 108 1010
nr. of particles per mm2 and 40 msec
incident radiation
object (photons)
2106
incident radiation
image intensifyer screen
2103
absorption efficiency
image intensifyer screen
103
light (photons)
image intensifyer screen
5106
electrons
image intensifyer screen
105
light (photons)
image intensifyer exit screen
109
Particle statistics in a typical X-ray chain

## Slide 14
Digital X-ray
Chapter 2.6-2.7

## Slide 15: Image conversion
Luminescence using barium fluorohalide (BaFBr:Eu)
incident radiation
absorbed x-rays liberate e- from valence band to conduction band; e- get trapped
scanning red laser stimulates trapped electrons
reflected light is collected in photo multiplier tube + stored as digital image
BaFBrEu
e-
e-
valence band
conduction band
PMT
red laser
1)
2)
3)
Notes: Notes
In semiconductors and insulators, there is a bandgap above the valence band, followed by a conduction band above that.
In metals, the conduction band has no energy gap separating it from the valence band. The rest of this article refers to the valence band in semiconductors and insulators.
PMT = Photo Multiplier Tube

## Slide 16: Digital Luminescence
incident radiation
exposure
excited state
ground state
remaining excitation
is removed by exposure
to intensive light
reset
read-out
light
laser beam
scan
stimulated luminescence
Ba F Br: Eu++
Notes: Notes
The method makes use of electrons trapped between the valence and electron band in the crystaline structure of matter. Aluminum oxide is one such type of matter. The trapping sites are imperfections of the lattice - impurities or defects. The ionizing radiation produces electron-hole pairs - electrons are in the conduction band and holes in the valence band. The electrons which have been excited to the conduction band may become trapped in traps present. Under stimulation of light the electrons may free themselves from the trap and get into the conduction band. From the conduction band they may recombine with holes trapped in hole traps. If the centre with the hole is a luminescence centre emission of light will occur. The photons are detected using a photomultiplier tube. The signal from the tube is then used to calculate the dose that the material had absorbed.

## Slide 17: Selenium Detector
exposure
read-out
charge
selenium
selenium
viewing storage imaging
(laser camera)
computer
Notes: Notes
Surface of selenium detector is charged
X-ray photons liberate electrons in the centre which neutralize surface charges
Surface potential is read

## Slide 18: Silicium Detector

## Slide 19: Scintillation Detector
Notes: Notes
A scintillator is a or substance that absorbs high energy (ionizing) electromagnetic or charged particle radiation then, in response, fluoresces photons at a characteristic Stokes-shifted (longer) wavelength, releasing the previously absorbed energy. See also Scintillation (physics). Scintillators are defined by their light output (number of emitted photons per unit absorbed energy), short fluorescence decay times, and optical transparency at wavelengths of their own specific emission energy. The latter two characteristics set them apart from phosphors. The lower the decay time of a scintillator, that is, the shorter the duration of its flashes of fluorescence are, the less so-called "dead time" the detector will have and the more ionizing events per unit of time it will be able to detect.
Scintillators are used in many physics research applications to detect electromagnetic waves or particles. There, a scintillator converts the energy to light of a wavelength which can be detected by inexpensive or easy to handle detectors such as photomultiplier tubes (PMTs).

## Slide 20: Contrast agent
Time
Absorption
Iodine contrast media to increase x-ray absorption
Chapter 2.9

## Slide 21: Contrast agent

## Slide 22: Digital Subtraction
pre-contrast
post-contrast
difference
Chapter 2.10
difference+enhancement

## Slide 23: DSA Angiography

## Slide 24: Real-time DSA Angiography

## Slide 25: Clicker Activity (5 min)
?
Digital Subtraction Angiography is performed including a pre-contrast measurement #1 and a post-contrast measurement #2. The contrast agent increases the linear attenuation coefficient in blood. What is the image contrast of projection through the center point of the setup?
• Contrast ~ (mb*d - m*D)
• Contrast ~ (mb *D - m *d)
• Contrast ~ (mb - m)*d

## Slide 26: Clicker Activity (Notes)
d
D
D
pre-contrast
post-contrast
blood vessel (mb)
m
#1
#2
m
- d

## Slide 27
Computed Tomography
Chapter 2.12

## Slide 28: Problem
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5

## Slide 29: Problem
I1
I2
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5
Notes: Blackboard
-> calculate contrast in a) projection imaging (=(1738-1734)/3472 = 0.01%) vs b) tomographic imaging (=(63-35)/98 = 28%)

## Slide 30: Projection versus Tomographic imaging
Projection imaging
= 3D structures are collapsed onto 2D image:
- details are obscured - low soft-tissue contrast - not quantitative
Tomographic imaging
= 3D structure is resolved:
- slices (“gr: Tomos = slice”)
- Computed image reconstruction
- Projection imaging and CT share same principles of x-ray generation, interaction, detection

## Slide 31: CT – Reconstruction problem
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5
1) Homogeneous object, monochromatic radiation
Chapter 1.10,2.14

## Slide 32: CT – Reconstruction problem
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5
2) Inhomogeneous object, monochromatic radiation

## Slide 33: CT – Reconstruction problem
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5
3) Inhomogeneous object, polychromatic radiation

## Slide 34: CT – History
Johann Radon, 1917
- mathematical foundation for tomographic image reconstruction
Allan Cormack, 1963
- first experiments with axial tomography
Godfrey Hounsfield, 1972
- first clinical prototype; manufactured at EMI Ltd., England
(Cormack and Hounsfield shared Nobel prize in 1979)
Notes: Notes
Johann Radon (December 16, 1887–May 25, 1956) was a mathematician born in Litoměřice in Bohemia (now Czech Republic). He wrote a doctoral dissertation on calculus of variations in 1910, at the University of Vienna.
Allan MacLeod Cormack (February 23, 1924 – May 7, 1998) was a South African-born American physicist who shared a part of the 1979 Nobel Prize in Physiology or Medicine for his work on the computerized axial tomography (CAT) scan.
Sir Godfrey Newbold Hounsfield CBE (28 August 1919 – 12 August 2004) was an English electrical engineer who shared the 1979 Nobel Prize for Physiology or Medicine with Allan McLeod Cormack for his part in developing the diagnostic technique of computerized axial tomography (CAT).His name is immortalised in the Hounsfield scale, a quantitative measure of radiodensity used in evaluating CAT scans. The scale is defined in Hounsfield units (symbol HF), running from air at -1000 HF, through water at 0 HF, and up to bone at +1000 HF.

## Slide 35: CT – Allan Cormack

## Slide 36: CT – Godfrey Hounsfield

## Slide 37: CT – Principle
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5

## Slide 38: CT – Image reconstruction
Direct inversion (solving N unknowns from at least N equations)
Iterative reconstruction
(image estimation, pseudo-projection, minimization of residuum)
1.
2.
3.
4.

## Slide 39: CT – Image reconstruction
Backprojection
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5

## Slide 40: CT – Image reconstruction
Backprojection
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5
1 projection
3 projections
N projections
Image

## Slide 41: CT – Image reconstruction
Backprojection – Point Spread Function (PSF)
Point object
Image
PSF is obtained if 1) radiusobject  0
2) number of projections  
3) number samples per projection  

## Slide 42: CT – Image reconstruction
Radon transform
x
y

Detector
y
x
r

“Sinogram”
xcos + ysin = 0
Radon transform = Image encoding principle in CT
Source
r
Notes: Blackboard
-> illustrate R[] for =[0,45,90deg]

## Slide 43: Clicker Activity (5 min)
?
Which object does the sinogram below correspond to?
• Object #1
• Object #2
• Object #3
2
φ
r
0
0
90°
180°
1
3
y
x

## Slide 44: CT – Image reconstruction
Inverse transform (I)
“Wanted”
Radon found analytical solution but it’s numerically impractical
“Use Fourier transform”
Notes: Blackboard
-> illustrate coordinate systems [x,y] -> [p,q] and [r,s] -> [u,v]

## Slide 45: CT – Image reconstruction
Inverse transform (II)
Reconstruction formula
Notes: Blackboard
-> p = u cos, q = u sin
-> dpdq = u du d

## Slide 46: CT – Image reconstruction
Inverse transform (III)
F{P (u)} is weighted by |u|
F{P (u)} |u| is inversely transformed
and summed over all angles  (=backprojection)
Notes: Blackboard
-> r = x cos + y sin
-> F{P(u,+)} = F{P(-u,)}

## Slide 47: CT – Image reconstruction
Convolution backprojection
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5

## Slide 48: CT – Image reconstruction
without pre-filtering
with pre-filtering
Adapted from Kalender WA, Computed Tomography, ISBN 3-89578-216-5

## Slide 49: CT – Image reconstruction
Convolution backprojection process
From Kalender WA, Computed Tomography, ISBN 3-89578-216-5

## Slide 50: CT – Pencil beam
X-ray
Source
Object
Detector

## Slide 51: CT – Fan beam

## Slide 52: CT – Parallel fan beam

## Slide 53: CT – Cone beam

## Slide 54: CT – Hounsfield unit

## Slide 55: “The most important slides”