## Slide 1: Biomedical Imaging X-Ray Imaging (I)
Sebastian Kozerke
Institute for Biomedical Engineering, University and ETH Zurich, Switzerland

## Slide 2: Lecture notes
Handouts + Exercises
https://moodle-app2.let.ethz.ch/course/view.php?id=11546
Book
Introduction to Medical Imaging
Cambridge University Press
Nadine Barrie Smith, Andrew Webb
www.library.ethz.ch  online access

## Slide 3: Exercises
Matlab based (in groups of two); problem solving
13-15 HG E7 Lecture
15-16 HG E7 Exercise (preparation)
Tue
Wed
Thu
Fri
Sat
Sun
Mon
Tue
Exercise #X
15-16 HG E7 Exercise (solution, discussion)
13-15 HG E19 Computer room
Mon <23:59 Submit Solution in Moodle

## Slide 4: Today’s Learning Objectives
Describe mechanisms of X-ray generation and X-ray spectrum
Relate interaction of X-rays in tissue to tissue properties
Illustrate basic instrumentation for planar radiography
Derive relation between modular transfer and spatial resolution
Implement basic X-ray projection setup (Exercise)

## Slide 5: Medical imaging
Gamma
X-ray
UV
10-16
10-14
10-12
10-10
10-8
10-6
10-4
10-2
100
102
IR
Micro
Radio
104
wave length [m]
Electromagnetic spectrum and attenuation in the body
Notes: Wavelength
Gamma: 100 pm
X-ray: 10 pm-10 nm
Light: 390-760 nm
Micro: 0.1-30 cm
Radio: 10 cm – 10 km

## Slide 6: Medical imaging
MRI
X-ray, CT
PET
102
Gamma
X-ray
UV
10-16
10-14
10-12
10-10
10-8
10-6
10-4
10-2
100
IR
Micro
Radio
104
wave length [m]

## Slide 7: Wilhelm Conrad Röntgen, 1845-1923
7

## Slide 8: Wilhelm Conrad Röntgen, 1845-1923
8
Notes: Notes
When the German physicist Wilhelm Conrad Röntgen did experiments on Cathode ray's with a Crookes tube, he discovered while working, that a plate of barium platinocyanide (fluorescent crystals) on a nearby table in his workroom glowed when he activated the tube. Even when he covered the tube with black material it kept glowing. Roentgen named this new phenomena X-Rays.
Wilhelm Conrad Röntgen (in English: "William Conrad Roentgen") (March 27, 1845 – February 10, 1923) was a German physicist, of the University of Würzburg, who, on November 8, 1895, produced and detected electromagnetic radiation in a wavelength range today known as x-rays or Röntgen Rays, an achievement that earned him the first Nobel Prize in Physics in 1901.
Röntgen was born in Lennep (now a part of Remscheid), Germany, to a clothmaker. His family moved to Apeldoorn in the Netherlands when he was three years old. He received his early education at the Institute of Martinus Herman van Doorn. He later attended Utrecht Technical School, from which he was expelled for producing a caricature of one of the teachers, a "crime" he claimed not to have committed.
In 1865, he tried to attend the University of Utrecht without having the necessary credentials required for a regular student. Hearing that he could enter the Federal Polytechnic Institute in Zurich, today the ETH Zurich, by passing its examinations, he began studies there as a student of mechanical engineering. In 1869, he graduated with a Ph.D. from the University of Zurich.
In the next experiments he used photographic material and made his first X-Ray picture, the hand of his wife Anna Bertha!

## Slide 9: X-ray image (1895)
Original x-ray image made by W.C. Röntgen (hand of his wife)
9

## Slide 10: X-rays in medical imaging
Radiography
    Thorax, bones etc.
Computed Tomography (CT)
    Bones and soft tissue
    Angiography
Angiography
    Using contrast agent
    Digital subtraction angiography (DSA)
Fluoroscopy
    Guidance of interventional procedures

## Slide 11
Generation of X-rays
Chapter 2.2-2.3

## Slide 12: Generation
fast electrons
target with heavy atoms
X-ray photons
Electrons interact with inner e- shells of target (X-rays are not radioactive!)
Notes: Blackboard
-> radiation consists of quanta
-> radiation energy is integer multiple of single quantum energy
-> single quantum energy depends on frequency -> Planck: E = h   = e  U (h = 6.6210-34 Js; 1 Js = 1 W/s2)
-> because of energy quantization radiation can be considered a particle stream (particle = photon)
-> E = m  c2 -> impulse of photon: p = m  c = h   / c (photons always travel at speed of light)
Notes
The electronvolt (eV) is the amount of kinetic energy gained by a single unbound electron when it passes
through an electrostatic potential difference of one volt, in vacuum. In other words, it is equal to one volt
(1 volt = 1 joule per coulomb) times the charge of a single electron (in coulombs). 1 eV = 1.610−19 J

## Slide 13: Generation – X-ray tube
Acceleration of electron in vacuum tube (UA = 30-200 kV, IA = 50-100 mA)
13
Notes: Blackboard
-> wanted: min = h  c / e  U (U = 100 kV)
-> min = 6.610-34 Js  3108 m/s / 1.610-19 C / 1105 V = 1.25  10-11 m = 12.5 pm (remember: 1V = 1J/C)
Notes
-> 1 ångström (Å) = 10–10 metres = 0.1 nm = 100 pm
->  of typical diagnostic X-rays: 10 pm-10 nm

## Slide 14: Generation – Spectrum
photon energy
K-lines of
wolfram
characteristic
x-ray
without filter
photon flux
anode voltage
Wolfram target (1 mm Al prefilter)
Bremsstrahlung  continuous spectrum
Characteristic radiation  peaks/lines

## Slide 15: Generation – Bremsstrahlung
Fast-moving electron decelerates when it swings around a heavy nucleus:
  electron interacts with electrons in outer shells
  energy depends on distance at which electron passes nuclei
  spectrum continuous 0 -> Emax
  interaction depends on Z value and E of electrons
electron
nucleus
fast
slower
X-ray photon
c
Notes: Notes
Bremsstrahlung (help·info), (from the German bremsen, to brake and Strahlung, radiation, thus, "braking radiation" or "deceleration radiation"), is electromagnetic radiation produced by the acceleration of a charged particle, such as an electron, when deflected by another charged particle, such as an atomic nucleus. The term is also used to refer to the process of producing the radiation. Bremsstrahlung has a continuous spectrum. The phenomenon was discovered by Nikola Tesla during high frequency research he conducted between 1888 and 1897.

## Slide 16: Generation – Bremsstrahlung
photon energy
K-lines of
wolfram
characteristic
x-ray
without filter
photon flux
anode voltage
Wolfram target (1 mm Al prefilter)
Bremsstrahlung  continuous spectrum
Characteristic radiation  peaks/lines

## Slide 17: Generation – Characteristic radiation
+
electron
K
L
+
K
L
photo electron
“hole”
+
K
L
Fast-moving electron interacts tightly with bound electrons in target:
  K-shell electron is ejected
  Electron energy Eelectron > Ebinding, K
  Discrete energy of emitted radiation E = Ebinding,K – Ebinding,L
1)
2)
3)
Notes: Blackboard
-> emitted photon energy: ΔE = EN-EK = h  
-> exercise: what is the maximum energy of X-rays from a tungsten target?
-> solution: maximum energy x-rays are generated if a zero-energy electron jumps onto K shell
-> binding energy of electrons Ebinding = 13.6 eV  Z2eff
-> Zw = 74 -> Emax = h   = 13.6 eV (74-2)2 = 70.5 keV (Zeff ~ (Z–2) for nuclei with large Z since electrons with l=0 shield nuclei)
-> min = h  c / Emax = 610-34 Js  3108 m/s / ( 70.5103/1.610-19 J) = 40 pm

## Slide 18: Generation – Characteristic radiation
photon energy
K-lines of
wolfram
characteristic
x-ray
without filter
photon flux
anode voltage
Wolfram target (1 mm Al prefilter)
Bremsstrahlung  continuous spectrum
Characteristic radiation  peaks/lines

## Slide 19: Generation – Materials and Electron energy
Energy spectrum (bremsstrahlung, photon
flux as function of energy) in case of various
target materials with pre-filter):
tungsten Z = 74
molybdenum Z = 42
copper Z = 29
Energy spectrum (bremsstrahlung, photon
flux as function of electron energy) for various accelerating fields:
a: Theoretical curve without
prefilter and self-absorption
photon flux
a
photon flux
a
100 kV
150 kV
200 kV

## Slide 20: Clicker Activity (5 min)
?
Imagine the anode material is changed from a thick to a very thin structure (consisting of a single atomic layer) - how would the spectrum shown change?
• Slope of spectrum gets steeper
• Emax is shifted to the right
• Spectrum disappears completely
• Constant flux up to Emax is seen.

## Slide 21: Generation – Focal Spot
Anode
Cathode

## Slide 22: Generation – X-ray tube

## Slide 23: Generation – Heat production
Power absorption:
black body:  = 1
Stefan-Boltzman law:
 = 5.7  10-12 Wcm-2K-4
Wien’s displacement law:
b = 2.9  10–3 mK
Notes: Blackboard
-> P = U  I = 100 kV  100 mA = 10 kW
-> 1% X-ray + 99% heat!
-> How much heat is radiated per second and cm2 if TAnode = 2000 K given:
- absorption efficiency  = 0.8
-> W =     T4 = 73 Wcm-2 = 73 Jcm-2s-1

## Slide 24: X-ray tube – Stationary and rotating anode

## Slide 25: X-ray tube – History
Notes: In the early times Crookes or Hittorf tubes consists of a view simple electrodes projecting into a glass bulb. With one of these tubes Wilhelm Röntgen made his discovery for which he was awarded for the first Nobel Prize. The classic cold X-Ray tube was born, a tube with two electrodes. The tube developed into a three electrode tube with a platinum target due to the high melting point of this metal. When in 1913 William Coolidge and Lilienfeld made there first hot Cathode X-Ray tube, they replaced the cold Cathode with a heated spiral filament Cathode and tungsten Anode. The angle of the Anode in his tube was set at 45º. X-Ray's could now be better controlled and were more reliable. There was only a practical problem, cooling the Anode, (which got extremely hot) was a huge problem due to it's small size. A new design was developed, a heavy copper base to the Anode, sometimes with water cooling, this conducted the heat away and therefore increased the capacity of the tube to withstand a high current. In 1924 Philips developed the Metalix tube, the glass tube was fitted in a metal housing and also incorporated the principle of line focus. The Anode face was changed at an angle of 19º in respect to the Cathode. Finally the rotating Anode tube was developed. It was first sold in 1929 by Philips and called the Rotalix tube. In the rotating Anode tube the target is a tungsten or wolfram disc that spins so that the focal spot of the Cathode ray is changing and the heat is dissipated. Soon after the discovery of Conrad Roentgen the production of these tubes started, not only in the German factories of Gundelach, Müller, Pressler and Siemens. But also by Philips in Holland and Cuthbert Andrews in England. In the US by Victor and general Electric.

## Slide 26
Attenuation of X-rays
Chapter 2.4-2.5

## Slide 27: Terminology
Transmission
    Photons passing through the body
Absorption
    Partial or total absorption of photon energy
Scatter
    Photons diverted in a new direction
Attenuation
    Absorption + scatter

## Slide 28: Attenuation
Attenuation in body with transmission of 80%:
Object
Object
Object
100%
80%
64%
51%
Photo effect
Compton effect
Interaction with bound electrons

## Slide 29: Attenuation – Photo effect
+
K
L
+
K
L
photo electron
“hole”
+
K
L
1)
Photo effect leads to total absorption of photon:
  Generation of characteristic x-ray radiation with low energy
  Photo electrons cause biological “damage”  “ionizing radiation”
  Probability ~   Z3/E3
2)
3)
incident photon
Notes: Blackboard
-> energy of incident photon EIncidence = EEjection + ½  me  2

## Slide 30: Attenuation – Compton effect
Back-scatter
Side scatter
Forward scatter
incident photon
recoil electron
Compton effect leads to photon “bouncing”:
  Photon loses (part of) its energy to kinetic energy of electron
  Detection of photon changes
  Recoil electron responsible for “biological” damage
  Probability ~ /E
scattered photon
Notes: Blackboard
-> Compare with billiard hit -> sketch photon hitting an electron a) before hit b) after hit (after hit impulse of photon is p’, of electron P’e)
-> conservation of energy: p  c + mc2 = p’  c + E’e -> ( p – p’ + mc)2 = ( E’e/c )2
-> conservation of momentum: p – p’ = p’e
-> square both sides: p2 – 2pp’+p’2 = p’e2 (p  p’ = p  p’  cos )
-> further calc
-> ’- = h/(mc)  (1-cos ) (Compton effect)
-> Scatter problem!

## Slide 31: Attenuation – Mass absorption coefficient (’)
a Photoeffect
b Photoeffect (/), Comptoneffect (/), pair production (/)
water
water
lead
photon energy
photon energy
cm2/g
cm2/g
10 keV 100 keV 1 MeV 10 MeV
10 keV 100 keV 1 MeV 10 MeV 100 MeV
100
10
1
0.1
0.01
0.001
10
1
0.1
0.01

## Slide 32: Attenuation coefficient
atom
cross section 
x-ray
I(x)
I(x+dx)
Mass attenuation coefficient ’ [cm2/g]
Linear attenuation coefficient  = ’   [cm-1]
dx

Notes: Blackboard
-> mass absorption coefficient ’ =   NA / A (NA: Avogadro’s number = 6  1023 / mol, A: atomic weight)
-> sketch: shooting target

## Slide 33: Attenuation coefficient
I(x)
I(x+dx)
Linear attenuation coefficient 
dx

Beer-Lambert’s law
Notes: Blackboard
-> Beer-Lambert: I(x+dx)-I(x) = -Idx
-> dI/I = -dx -> ln|I| = -x + c -> I = I0 exp(-x)

## Slide 34: Attenuation – Linear attenuation coefficients (’)
bone
muscle
fat
μ
E [keV]
μ [cm-1]

## Slide 35: Attenuation – Summary
Photo effect 
    X-ray photon disappears
    Photo electron recoils
    Effective in contrast agent, lead, bone
Compton effect 
    Interaction with free electrons
    Scatter
    Water, air, soft tissue, bone
(Pair production  )
    Electron-positron pair requires E > 1 MeV
    Not relevant in medical imaging

## Slide 36: Clicker Activity (5 min)
?
Assign energy values to the two projection images and explain the reasons for the differences in image contrast and signal intensity
• left/right: same energy but different display window/level settings
• left: high energy, right: low energy
• left: low energy, right: high energy

## Slide 37
Projection imaging

## Slide 38: Projection imaging – Pre-filtering
Detector
Aluminum filter
Aluminum filter thickness
x-rays
Notes: Notes
-> Low energy photons do not reach the detector
- absorbed in x-ray tube material and patient (Dose)
- No contribution to the image
-> Inherent filtering in tube (glass, oil, target)
-> Additional filtering to reduce patient dose (aluminum sheets between

## Slide 39: Resolution of imaging procedures
Spatial resolution
    Blurring of object points when projected into image
Contrast
    Difference in linear attenuation coefficients between tissues
Temporal resolution
    Sample rate of temporal recording
Chapter 1.3

## Slide 40: Resolution – Spatial resolution
Object representation

## Slide 41: Resolution – Spatial resolution
Mapping object space onto image space
Object space
Image space
Object space
Image space
“Line Spread Function“ (LSF)
Notes: Blackboard
-> Calculate LSF for MRI
- LSF(x) = int(exp(jkx)) kmax=-kmin=/Δx
- LSF(x) = 2sin(x/Δx)/x
-> LSF in X-ray is determined experimentally with lead grids

## Slide 42: Resolution – Spatial resolution
Mapping object space onto image space (1D  2D)
1D: LSF  2D: PSF
PSF: Point Spread Function

## Slide 43: Resolution – Spatial resolution
Fourier series representation of object
Wave number (see:. spatial frequency = k-space in MRI):
Mapping Fourier representation of object onto image:
Notes: Blackboard
-> use x’’ = x’ – X
-> sin(X) = sin(x’-x’’) = sin(x’)cos(x’’) – cos(x’)sin(x’’)

## Slide 44: Resolution – Spatial resolution
Mapping Fourier representation of object onto image:
Object space
Image space
“MTF: Modulation Transfer Function”
Notes: Blackboard
-> use x’’ = x’ – X
-> sin(X) = sin(x’-x’’) = sin(x’)cos(x’’) – cos(x’)sin(x’’)

## Slide 45: Resolution – Modular Transfer Functions (MTF)
1
2
3
4
5
MTF in
lp/mm
25 cm image intensifyer,
CsJ screen
25 cm image intensifyer,
15 cm setting
17 cm image intensifyer,
CsJ screen
17 cm image intensifyer,
ZnS or CdS screen
Film- intensifying screen
combination
1:
2:
3:
4:
5:
lp/mm: line pairs per mm

## Slide 46: Clicker Activity (5 min)
?
Compare the Modular Transfer Function (MTF) of two different imaging systems
• The imaging system with MTF #1 yields best resolution
• The imaging system with MTF #2 yields best resolution
• Resolution of #1 and #2 is identical

## Slide 47: “The most important slides”