## Slide 1: Biomedical Imaging Nuclear Imaging (II)
Sebastian Kozerke / Markus Rudin
Institute for Biomedical Engineering, University and ETH Zurich, Switzerland

## Slide 2: Today’s Learning Objectives
Explain various PET coincidence events and corrections thereof
Describe SPECT radioisotope production (Technetium generator)
Describe PET radioisotope production (cyclotron)
Derive basic kinetic models for PET tracer quantification
Implement kinetic modeling for PET (Exercise)
2
2

## Slide 3: Positron Emission Tomography
e-
g(511 keV)
g(511 keV)
e+
Positron
emitter
Notes: Registration of event:
Coincidence detection, i.e.
two events occuring within a
time window Dt attributed to
One anihilation event
| t1 - t2 | < Dt
Resolution:
- Positron lifetime -> 0.3-1 mm
Number of detectors and size of detector
Off-center position
Setting of t
colineraity

## Slide 4: Activity (5 min)
Using the mass of positron/electron, show that the energy of g-photons is indeed 511 keV (mp,e = 9.11∙10-31 kg; 1 eV = 1.6 ∙10-19 J)!
e-
g(511 keV)
g(511 keV)
e+
Positron
emitter

## Slide 5: Coincidence types
true coincidence
scattered coincidence
random coincidence
multiple coincidence
Chapter 3.18

## Slide 6: Coincidence detection
A
B
A
B
A+B
Coincidence window 2t
t
threshold

## Slide 7: Attenuation correction
xij
Dij-xij
Loss of g-photons due to Compton scattering
Total scattering probability
Chapter 3.18.1

## Slide 8: Attenuation correction
Dij
m
i
j
internal PET source
Dij
m
i
j
=
external PET source

## Slide 9: Coincidence types
true coincidence
scattered coincidence
random coincidence
multiple coincidence

## Slide 10: Scatter correction
Analytical method
assumption: scatter distribution varies slowly across scattering medium
use phantom data to correct for scattered events
Dual energy method
Energy range 400-600 keV: scattered and true events
Energy range 200-400 keV: only scattered events
Simulation method
Reconstruction without scattering correction
Estimation of scattering correction
Chapter 3.18.3

## Slide 11: Coincidence types
true coincidence
scattered coincidence
random coincidence
multiple coincidence

## Slide 12: Random coincidences
Frequency of random events
Individual photon detection rates of a pair of detectors:
Coincidence window:
Frequency of random events:
‚Randoms‘
no geometrical information
almost uniform distribution across field-of-view
 reduced contrast-to-noise ratio
randoms
Chapter 3.18.2

## Slide 13: Correction of random coincidences
Estimate number of ,randoms’ for each detector pair: N1, N2, 2t must be known
Measure/estimate number of ‚randoms‘ by adding parallel coincidence circuit
Coincidence window 1
Coincidence window 2
A
B
A
B
‚trues‘ and ‚randoms‘
‚randoms‘
Estimation of number of true events:

## Slide 14: Coincidence types
true coincidence
scattered coincidence
random coincidence
multiple coincidence
See random coincidence

## Slide 15: Dead time correction
Detector dead time: number of events measured < number of true events
Effect more severe for high activities  high count rate, high fraction of missed events:
ideal ‚trues‘
actual ‚trues‘
d…dead time
d = 1ms
Chapter 3.18.4

## Slide 16: 3D image reconstruction
Data are reduced to a set of parallel 2D sinograms by rebinning all sinograms with same z value into the same sinogram
z1
z2
z4
z3
z5
z6
z7
z8
▶︎ Reconstruction of N 2D sinograms
Chapter 3.16

## Slide 17: 3D image reconstruction

## Slide 18: Image noise reduction
see e.g. exercise XCT #2
Chapter 3.19

## Slide 19: Image noise reduction

## Slide 20: Clinical PET scanner

## Slide 21: Animal PET scanner

## Slide 22: PET/CT scanner

## Slide 23: Total body PET scanner

## Slide 24
Radiotracer Production

## Slide 25: SPECT radioisotopes
Radioisotope | Energy (keV) | Half-life | Production
67Ga | 93, 185, 300 | 3.3d | cyclotron
99mTc | 140 | 6h | generator
111In | 173, 247 | 67h | cyclotron
123I | 160 | 13h | cyclotron
133Xe | 81 | 5.2d | reactor
201Tl | 60, 83 | 73h | cyclotron
Chapter 3.3
Notes: TI thalium

## Slide 26: Technetium generator
Eluate: NaCl
Lead screen
Generator:
Mother nuclide (99Mo)
Immobilized on alumina
(Al2O3)
Collection vial for
daughter nuclide
(99mTc)
Lead screen
Chapter 3.4

## Slide 27: Technetium generator
(67 h) Mo
99
42
(6 h) Tc
99m
43
(2105 y) Tc
99
43
Ru
99
44
-
82%
142 keV
140 keV
1% 99%
-
g (140 keV)

## Slide 28: Technetium generator
kMo
kTc
2.9010-6s-1
3.2010-5s-1
Kinetic equations:
Solution:
Radioactivity of 99mTc:

## Slide 29: Technetium generator
Time (h)

## Slide 30: Technetium tracers
Radiolabel | Ligand | Application
99mTcO4-Na | Sodium pertechnate | Thyroid function
99mTc-HSA | Human serum albumine | Blood pool, cardiac function
99mTc-HSA | Human serum albumine (microspheres) | Regional perfusion, vascular patency
99mTc-MIBI | 2-methoxy-isobutylisonitril | Myocardial function
99mTc-HMPAO | Hexamethyl-propylenaminoxim | Cerebral perfusion
99mTc-HEDP | Hydroxyl-ethylene-diphosphonate | Bone metabolism / formation
99mTc-MDP | Methylene-diphosphonate | Bone metabolism / formation
99mTc-DTPA | Diethylene-triamine-tetraacetate | Kidney function, lung function

## Slide 31: SPECT imaging of bone

## Slide 32: PET radioisotopes
Radionuclide | Half-life | Emax (MeV) | e+ decay (%)
11C | 20.4 min | 0.96 | 100
13N | 10 min | 1.20 | 100
15O | 2 min | 1.73 | 100
18F | 110 min | 0.63 | 97
22Na | 2.6 y | 0.55 | 90
62Cu | 9.7 min | 2.93 | 97
64Cu | 12.7 h | 0.65 | 29
68Ga | 67.6 min | 1.89 | 89
76Br | 16.2 h | Various | 56
82Rb | 1.3 min | 2.60, 3.38 | 96
124I | 4.2 d | 1.53, 2.14 | 23
Chapter 3.14

## Slide 33: Production of 18F radionuclide
https://www.youtube.com/watch?v=zD1fnNAmAPk

## Slide 34: Properties of 18F radionuclides
Lifetime 110 min
Lifetime reasonably long for chemical synthesis
Lifetime long enough to allow for shipping of labeled compound
18F PET does not require on-site cyclotron
Studies can be repeated in relatively short time intervals
but
Pharmacokinetics different from native, endogenous compound

## Slide 35: PET radioisotopes
Radionuclide | Half-life | Emax (MeV) | e+ decay (%)
11C | 20.4 min | 0.96 | 100
13N | 10 min | 1.20 | 100
15O | 2 min | 1.73 | 100
18F | 110 min | 0.63 | 97
22Na | 2.6 y | 0.55 | 90
62Cu | 9.7 min | 2.93 | 97
64Cu | 12.7 h | 0.65 | 29
68Ga | 67.6 min | 1.89 | 89
76Br | 16.2 h | Various | 56
82Rb | 1.3 min | 2.60, 3.38 | 96
124I | 4.2 d | 1.53, 2.14 | 23
Chapter 3.14

## Slide 36: Production of 11C radionuclide
Proton
capture
Proton source:
cyclotron
Metastable
nucleus
(t1/2 << 1s)
Stabilization by emission of a-particle
Positron
emitter
Stabilization by
positron emission
Chemical
synthesis
Stable
nucleus
14N + p [15F ]
(7p, 7n) (8p, 7n)
a
(2p, 2n)
11B
(5p, 6n)
11C
(6p, 5n)
e+ + n
20 min
<< 1s

## Slide 37: Properties of 11C radionuclides
All organic compounds contain C and can be labeled in principle
Pharmacokinetic properties identical to that of ‚cold‘ parent compound
Studies of drug biodistribution / PK
Receptor affinity identical to that of parent compound
but
Short half life (20 min)
Chemical synthesis difficult due to time constraints

## Slide 38
Quantitative PET Analysis

## Slide 39: Quantitative PET – definitions
Injected dose per gram of tissue:
Standardized Uptake Value (SUV):
M : Body mass
ct tissue concentration
vt volume of tissue ROI
mt mass of tissue ROI
Dinj injected dose

## Slide 40: Quantitative PET – receptor binding
Ligand
Cell
Receptor
Cell
PET radiotracer
Ligand
Receptor

## Slide 41: Quantitative PET – receptor binding
Receptor Binding: L + R  RL Equilibrium constant: (principle of microreversibility)
Mass conservation for receptor:
yields
or the receptor binding
and the Scatchard equation
(L: Ligand, R: Receptor)

## Slide 42: Quantitative PET – receptor binding
Binding plot
Scatchard plot
for [RL]<<[RT]
[L]
log[L]
Kd
[RL]/[Rt]
[RL]/[Rt]
Bmax: maximum binding capacity

## Slide 43: Quantitative PET – tracer extraction
Renkin-Crone equation
The amount of tracer extracted from capillaries depends on capillary surface S, the vascular permeability P and the blood flow F:
F
F
ca
cv
Tracer extraction from capillaries

## Slide 44: Quantitative PET – volume distribution
cb
ct
v1
vt
cb
ct
tissue
tissue
blood
blood
Distribution volume
Tracer experiment
M: total amount of tracer

## Slide 45: Clicker Activity (5 min)
After administration of 1 μg/kg of a tracer into a patient of 70 kg body mass, an average blood concentration of 1.4 mg/l is measured. What is the distribution volume?
  • 5 l
  • 500 ml
  • 50 l
  • 10 l

## Slide 46: Quantitative PET – compartment modeling
Blood
Target organ
‚tissue-reservoirs‘

## Slide 47: Quantitative PET – modeling of receptor binding
Tracer fraction in tissue
compartments
p: plasma
f: free (in tissue)
ns: non-specific bound
b: bound
PET activity at time t:
i = p, f, ns, b
vi fractional volume of compartment

## Slide 48: Quantitative PET – modeling of receptor binding
Equilibrium free / non-specifically bound tracer instantaneous:
Kinetic equations:

with rate constants:
Renkin-Crone:

## Slide 49: Quantitative PET – tracer kinetics for receptor binding
Solution for compartment expressing receptor:
with coefficients
Solution for compartment not expressing receptor:

## Slide 50: Quantitative PET – tracer binding example
ct(t)
k1=1min-1, k2=0.1min-1, k3=1min-1 and 0.5min-1, k4=0.01min-1
Compartment expressing receptor
Compartment not expressing receptor

## Slide 51: Quantitative PET – cell energy metabolism

## Slide 52: Quantitative PET – cell energy metabolism
CH2OPO32
CH2OH
OH
HO
HO
O
1
2
3
4
5
6
HO
HO
OH
OH
CH2OH
O
1
2
3
4
5
6
HO
HO
OH
OH
CH2OPO32-
O
1
2
3
4
5
6
HO
HO
18F
OH
CH2OH
O
1
2
3
4
5
6
HO
HO
OH
CH2OPO32-
O
1
2
3
4
5
6
X
18F
hexokinase
isomerase
Glucose (G)
[18F]-2-fluoro-2-
deoxyglucose (FDG)
Glucose -6-phosphate(G6P)
FDG6P
Fructose-6-phosphate(G6P)
Ligand trapped
in cells as FDG6P
Native glucose utilization
FDG PET deoxyglucose utilization

## Slide 53: Quantitative PET – FDG compartment models
-
-

## Slide 54: Quantitative PET – FDG compartment models
Total tracer concentration in tissue
(assume vp<<vtV)
Kinetic equations for tracer:
Solution:

## Slide 55: Quantitative PET – FDG compartment models
Metabolic rate of FDG:
Metabolic rate of glucose (glucose utilization):
Lumped constant (ca. 0.6 in humans)
Correction factor to translate measured reaction rate for
FDG to that for glucose

## Slide 56: Quantitative PET – FDG lumped constant
Patient with brain glioma. The panels include calculated metabolic rate for the true tracer [1-11C]glucose and the analog FDG. The two PET images were divided (PETFDG/PETglc) on a pixel-by-pixel basis to provide an image of regional variation of the lumped constant LC.
Krohn et al (2005) Nucl Med Biol. 32:663

## Slide 57: Application – FDG PET of Alzheimer disease
Reduced FDG PET activity in Alzheimer patient compared to age-matched control

## Slide 58: Application – FDG PET of cocaine abuse
Cocaine abuse
Normal

## Slide 59: Application – FDG PET of cancer metastases

## Slide 60: “The most important slides”

## Slide 61: Quantitative PET – tracer kinetics for receptor binding
Solution applying basis transformation c  y
Coupled differential equations:
Decoupling by searching eigenvectors:
Eigenvalues:

## Slide 62: Quantitative PET – competition cold / hot ligand
‚hot‘ ligand
‚cold‘ ligand
Modified tracer kinetics in the presence of a cold drug:
Two ligands (drug, tracer) competing for the same receptor
Kinetic equations for

## Slide 63: Quantitative PET – competition cold / hot ligand
[RL*] + [RL] + [R] = [Rt]
Receptor occupancy for cold drug:
Competition (drug injected prior to tracer experiment)
Displacement (drug injected during tracer experiment)
The two equation systems are coupled through the mass balance of the receptor
* hot ligand
or
B’max
Bmax
Bmax
B’max

## Slide 64: Quantitative PET – competition cold / hot ligand
Bmax(D2)
B’max(D2;H)
B’max(C)
B’max(D2;R)
B’max(D2;C)
B’max(D2;O)
Bmax(5HT2)
B’max(5HT2;H)
B’max(5HT2;R)
B’max(5HT2;C)
B’max(5HT2;O)
Notes: Schizophrenic patient under various antipsychotic drugs. Interaction with dopamine D2 receptors

## Slide 65: Flortaucipir PET – detection of brain plaques
https://www.nejm.org/doi/full/10.1056/NEJMoa1900757
Notes: PET maps superimposed on MRI brain scans show voxels with higher regional:cerebellar gray-matter flortaucipir standard uptake value ratios (SUVRs) among former National Football League (NFL) players.
We used flortaucipir positron-emission tomography (PET) and florbetapir PET to measure deposition of tau and amyloid-beta, respectively, in the brains of former National Football League (NFL) players with cognitive and neuropsychiatric symptoms and in asymptomatic men with no history of traumatic brain injury. Automated image-analysis algorithms were used to compare the regional tau standardized uptake value ratio (SUVR, the ratio of radioactivity in a cerebral region to that in the cerebellum as a reference) between the two groups and to explore the associations of SUVR with symptom severity and with years of football play in the former-player group.